## valve box-jumper
Part | Name | Equipment | Action/Location | Size/model | Range | I/O | Signal
Vacuum | V001 | Vacuum vessel | valve box |  |  |  |  | tbd | to be defined
equipment and | V002 | Vacuum vessel | jumper |  |  |  |  | AO | Analog output
room temperature | CV001 | Control valve | pressurized GHe to 40K circuit | tbd | tbd | AO | 4-20mA | AI | Analog input
circuits | CV002 | Control valve | pressurized GHe to 5K circuit | tbd | tbd | AO | 4-20mA | DO | Digital ouput
| FV001 | Stop valve | 40K circuit to pumping line | DN25 | tbd | DO |  | DI | Digital input
| FV002 | Stop valve | 4.5K circuit to pumping line | DN25 | tbd | DO
| HV001 | Manual valve | isolation valve, PT001 & PT002 | DN25 | tbd
| HV002 | Manual valve | pumping valve, PT001 & PT002 | DN25 | tbd
| HV006 | Manual valve | isolation valve | DN40 | tbd
| PT001 | Secondary gauge | valve box + jumper | DN25 | tbd | AI
| PT002 | Primary gauge | valve box + jumper | DN25 | tbd | AI
| PV001 | Helium Pressure regulator | helium gas to purge the helium circuits |  | tbd
| SV001 | Safety valve | vacuum vessel cryogenic line | tbd | tbd
| SV002 | Safety valve | vacuum vessel valve box | tbd | 0.02barg
Helium - cold circuits | HX300 | Subcooling Heat exchanger
| CV100 | Cryogenic control valve | 40K circuit inlet from TSI circuit | tbd | tbd | AO | 4-20mA
| CV101 | Cryogenic control valve | 40K circuit outlet to TSO circuit | tbd | tbd | AO | 4-20mA
| CV200 | Cryogenic control valve | 4.5K circuit inlet from HP circuit | tbd | tbd | AO | 4-20mA
| CV201 | Cryogenic control valve | 4.5K circuit outlet to AP circuit | tbd | tbd | AO | 4-20mA
| CV202 | Cryogenic control valve | 4.5K circuit inlet to cavities cool down circuit | tbd | tbd | AO | 4-20mA
| CV300 | Cryogenic control valve | 2K circuit outlet to AP circuit | tbd | tbd | AO | 4-20mA
| CV301 | Cryogenic control valve | 2K circuit outlet to LP circuit | tbd | tbd | AO | 4-20mA
| CV302 | Cryogenic control valve | 2K circuit inlet from HX300 outlet | tbd | tbd | AO | 4-20mA
| FV100 | Cold stop valve | 40K circuit outlet to recovery line | tbd | tbd | DO
| FV200 | Cold stop valve | 4.5K circuit outlet to recovery line | tbd | tbd | DO
| PT100 | Pressure sensor | 40K circuit | tbd | 0–20bara | AI
| PT200 | Pressure sensor | 4.5K circuit | tbd | 0–7bara | AI
| SV100 | Safety valve | Thermal shield circuit | tbd | tbd
| SV200 | Safety valve | 4.5K circuit | tbd | tbd
| TT100 | Pt thermometer | 40K circuit inlet | Pt100 Wire-Wound | 35K-330K | AI | Pt100 card
| TT101 | Pt thermometer | Thermal shield | Pt100 Thin film | 35K-330K | AI | Pt100 card
| TT102 | Pt thermometer | Thermal shield | Pt100 Thin film | 35K-330K | AI | Pt100 card
| TT103 | Pt thermometer | 40K circuit outlet | Pt100 Wire-Wound | 35K-330K | AI | Pt100 card
| TT200 | Cernox thermometer | 4.5K circuit inlet | CX-1050-CD | 1.5K-330K | controller | Ethernet
| TT201 | Cernox thermometer | 4.5K circuit to HX300 inlet | CX-1050-CD | 1.5K-330K | controller | Ethernet
| TT202 | Cernox thermometer | HX300 outlet to CV302 inlet | CX-1050-CD | 1.5K-330K | controller | Ethernet
| TT203 | Cernox thermometer | 4.5K circuit outlet to AP circuit | CX-1050-CD | 1.5K-330K | controller | Ethernet
| TT300 | Cernox thermometer | HX300 outlet to 2K circuit outlet | CX-1050-CD | 1.5K-330K | controller | Ethernet
| TT301 | Cernox thermometer | 2K phase separator to HX300 inlet | CX-1050-CD | 1.5K-330K | controller | Ethernet

## cryomodule
Part | Name | Equipment | Action/Location | Size/model | Range | I/O | Signal
Vacuum | V003 | Vacuum vessel | cryomodule |  |  |  |  | tbd | to be defined
equipment and | J001 | Vacuum pumping system | cryomodule |  |  |  |  | AO | Analog output
room temperature | CV003 | Control valve | cavity 1 cooling water outlet | tbd | tbd | AO | 4-20mA | AI | Analog input
circuits | CV004 | Control valve | cavity 2 cooling water outlet | tbd | tbd | AO | 4-20mA | DO | Digital output
| EH001 | Electrical heater | cavity 1 coupler window | tbd | tbd | AO | 4-20mA | DI | Digital input
| EH002 | Electrical heater | cavity 2 coupler window | tbd | tbd | AO | 4-20mA | Lact | Active length
| FT001 | Flowmeter | cavity 1 cooling water outlet | tbd | tbd | AI |  | Ltot | Total length
| FT002 | Flowmeter | cavity 2 cooling water outlet | tbd | tbd | AI
| FV003 | Stop valve | 2K circuit outlet to pumping line | DN25/40 |  | DO
| FV004 | Stop valve | vacuum vessel to pumping system | DN150 |  | DO
| FV005 | Gate valve | cavities/beam vacuum | tbd | tbd | DO
| FV006 | Gate valve | cavities/beam vacuum | tbd | tbd | DO
| HV003 | Manual valve | isolation valve, PT003 & PT004 | DN25
| HV004 | Manual valve | pumping valve, PT003 & PT004 | DN25
| HV005 | Manual valve | isolation valve | DN40
| PT003 | Primary gauge | cryomodule | DN25 |  | AI
| PT004 | Secondary gauge | cryomodule | DN25 |  | AI
| PT005 | Pressure sensor | cooling water | tbd | 0-10bara | AI
| SV003 | Safety valve | vaccum vessel | tbd | 0.02barg
| TT001 | Pt thermometer | cavity 1 coupler window |  |  | AI | Pt100 card
| TT002 | Pt thermometer | cavity 2 coupler window |  |  | AI | Pt100 card
| TT003 | Pt thermometer | cavity 1 cooling water inlet circuit |  |  | AI | Pt100 card
| TT004 | Pt thermometer | cavity 2 cooling water inlet circuit |  |  | AI | Pt100 card
| TT005 | Pt thermometer | cavity 1 cooling water outlet circuit |  |  | AI | Pt100 card
| TT006 | Pt thermometer | cavity 2 cooling water outlet circuit |  |  | AI | Pt100 card
Helium - cold circuits | EH300 | Electrical heater | cavity 1 warm up | ThermoFoil | 100W, 48V | AO | 4-20mA
| EH301 | Electrical heater | cavity 1 warm up cold tuner | ThermoFoil | 100W, 48V | AO | 4-20mA
| EH302 | Electrical heater | cavity 2 warm up | ThermoFoil | 100W, 48V | AO | 4-20mA
| EH303 | Electrical heater | cavity 2 warm up cold tuner | ThermoFoil | 100W, 48V | AO | 4-20mA
| FV300 | Cold stop valve | 2K circuit outlet to recovery line | tbd | tbd | DO
| LE300 | Liquid helium level probe | cavity helium tank | tbd | Lact:tbd, Ltot:tbd
| LI300 | Level indicator | level probe supply and measure |  |  | AI/Ethernet
| PT300 | Pressure sensor | atmospheric pressure range | tbd | 0-2bara | AI
| PT301 | Pressure sensor | low pressure range | tbd | 0-0.1bara | AI
| RD300 | Rupture disc | cavity helium tank | tbd | tbd
| RD301 | Rupture disc | cavity helium tank | tbd | tbd
| SV300 | Safety valve | 2K circuit | tbd | tbd
| TT104 | Pt thermometer | 40K circuit inlet | Pt100 Wire-Wound | 35K-330K | AI | Pt100 card
| TT105 | Pt thermometer | beam tube thermalization | Pt100 Thin film | 35K-330K | AI | Pt100 card
| TT106 | Pt thermometer | cavity 1 coupler thermalization | Pt100 Thin film | 35K-330K | AI | Pt100 card
| TT107 | Pt thermometer | cavity 2 coupler thermalization | Pt100 Thin film | 35K-330K | AI | Pt100 card
| TT108 | Pt thermometer | beam tube thermalization | Pt100 Thin film | 35K-330K | AI | Pt100 card
| TT109 | Pt thermometer | thermal shield top | Pt100 Thin film | 35K-330K | AI | Pt100 card
| TT110 | Pt thermometer | thermal shield bottom | Pt100 Thin film | 35K-330K | AI | Pt100 card
| TT111 | Pt thermometer | 40K circuit outlet | Pt100 Wire-Wound | 35K-330K | AI | Pt100 card
| TT204 | Cernox thermometer | cavity 1 coupler thermalization | CX-1050-CD | 1.5K-330K | controller | Ethernet
| TT205 | Cernox thermometer | cavity 1 coupler thermalization | CX-1050-CD | 1.5K-330K | controller | Ethernet
| TT206 | Cernox thermometer | 4.5K circuit inlet to cool down cavities | CX-1050-CD | 1.5K-330K | controller | Ethernet
| TT207 | Cernox thermometer | Inlet to cool down cavities | CX-1050-CD | 1.5K-330K | controller | Ethernet
| TT302 | Cernox thermometer | CV302 outlet to 2K phase separator | CX-1050-CD | 1.5K-330K | controller | Ethernet
| TT303 | Cernox thermometer | 2K phase separator to HX300 inlet | CX-1050-CD | 1.5K-330K | controller | Ethernet
| TT304 | Cernox thermometer | 2K phase separator top | CX-1050-CD | 1.5K-330K | controller | Ethernet
| TT305 | Cernox thermometer | cavity 1 helium tank top | CX-1050-CD | 1.5K-330K | controller | Ethernet
| TT306 | Cernox thermometer | cavity 1 electrical heater | CX-1050-CD | 1.5K-330K | controller | Ethernet
| TT307 | Cernox thermometer | cavity 1 coupler | CX-1050-CD | 1.5K-330K | controller | Ethernet
| TT308 | Cernox thermometer | cavity 2 helium tank top | CX-1050-CD | 1.5K-330K | controller | Ethernet
| TT309 | Cernox thermometer | cavity 2 electrical heater | CX-1050-CD | 1.5K-330K | controller | Ethernet
| TT310 | Cernox thermometer | cavity 2 coupler | CX-1050-CD | 1.5K-330K | controller | Ethernet